## ProbeGainMeasurements
Test Setup:	Revision 2 Board	20mVPP from Keysight 33210A in high impedance output mode	Keysight Ouptut divided by resistor divider 22k and 1k	Signal Swept from 10Hz-20KHz	Output read from Keysight 2022A FFT in dB

Board 4	Now Dead. 823 is sinking current through negative rail, Something wrong with an internal amplifier it seems like, it doesn’t have the same ohm measurement between out pin and gain pin		Board 3			Board 2			Rev 1 Board 1				Rev 1 Board 2			Board 7	Using 11.1V batteries			Board 6	Using 11.1V batteries		Board 5	Using 11.1V batteries		Board 3 03/27/24	Using 11.1V batteries

Frequency	Magnitude (dBV)		Frequency	Magnitude (dBV)		Frequency	Magnitude (dBV)		Frequency	Magnitude (dBV)	Magnitude (dB)		Frequency	Magnitude (dBV)		Frequency	Magnitude (dBV)	Magnitude (dB)		Frequency	Magnitude (dBV)		Frequency	Magnitude (dBV)		Frequency	Magnitude (dBV)
110	-9.2		150	-7.1		110	-6.8		110	-7.5	62.7448566769917		110	-7.5		100	-7.1	63.1448566769917		100	-7.2		100	-7.1		100	-7.1
150	-9.5		190	-7.1		150	-7		150	-7.7	62.5448566769917		150	-7.7		122	-7.1	63.1448566769917		259	-7.1		229	-7		427	-7.1
230	-9.3		310	-7.1		340	-7		230	-7.6	62.6448566769917		230	-7.5		244	-7	63.2448566769917		633	-7.2		389	-7.1		710	-7.3
340	-9.5		420	-7.1		690	-7.2		340	-7.8	62.4448566769917		340	-7.8		580	-7.1	63.1448566769917		1404	-7.2		786	-7.4		984	-7.6
460	-9.6		610	-7.3		990	-7.5		690	-7.9	62.3448566769917		460	-7.8		1579	-7.2	63.0448566769917		2785	-7		1480	-8.1		1434	-8.1
950	-10.2		880	-7.5		1410	-8.1		990	-8.4	61.8448566769917		690	-8		3326	-6.9	63.3448566769917		3998	-6.8		2602	-9.6		1862	-8.8
1410	-10.7		1260	-8		2020	-9.1		1410	-8.8	61.4448566769917		950	-8.2		5142	-6.7	63.5448566769917		4746	-6.7		4105	-11.7		2228	-9.3
2020	-11.6		1790	-8.9		2940	-10.5		2020	-9.6	60.6448566769917		1410	-8.8		8530	-6.9	63.3448566769917		5837	-6.6		5676	-13.6		3609	-11.3
4120	-14.8		2590	-10.2		4200	-12.3		2900	-11.1	59.1448566769917		1980	-9.8		11795	-7.7	62.5448566769917		7202	-6.7		8324	-16.5		6065	-14.5
5910	-17.4		5340	-14.3		6030	-14.8		4160	-13	57.2448566769917		2900	-11.1		13733	-8.3	61.9448566769917		9056	-7		9621	-17.9		8118	-16.8
8540	-20.1		7710	-17.1		8660	-17.6		5990	-15.4	54.8448566769917		4160	-13		16556	-9.2	61.0448566769917		12489	-7.9		13275	-21.3		10544	-19.3
12250	-23.1		11100	-20.2		12440	-20.5		8620	-18.4	51.8448566769917		5950	-15.4		18219	-9.7	60.5448566769917		14046	-8.3		16594	-23.9		12222	-20.9
			15910	-23.1		17850	-23.6		12400	-21.2	49.0448566769917		8580	-18.3		19402	-10.1	60.1448566769917		15778	-8.8		19234	-25.7		15450	-23.5
			19680	-25					17780	-24.4	45.8448566769917		12320	-21.3				70.2448566769917		17700	-9.5					18806	-26
										17.82	-24.5
Output Comparison of Separate Probe Boards
Frequency
Magnitude (dBV)
Output Comparison of Separate Probe Boards
Frequency
Magnitude (dBV)
Gain Comparison of Multiple Probe Boards and Revisions
Frequency (Hz)
Magnitude (dB)